## RDS
Standard Components - Room Data Sheet
Date of issue: 05.04.2018
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
BIRM-A | Birthing Room - LDR Without Bath | 34 | 24 Hours | 1 mother; 1-2 support persons; up to 4 staff |  | Sound attenuation level high, Acoustic privacy is required. | The Birthing Room LDR (without birthing bath/ pool) provides for the entire process of delivery including mother's preparation, relaxation and analgesia during labour, delivery, infant resuscitation, post natal recovery and observation. The Birthing Room will include a scrub basin. Clinical items such as medical gases and services may be concealed but within easy reach. The room will have direct access to an Ensuite - Birthing. The room can be used for an extended stay where LDPR models are used. Note that Room Area includes 3m2 Equipment Store. Notes applied to floor ratings respond to current AS 4586 and associated HB198 requirements - compliance with current regulations and jurisdictional policy must be verified by project teams on each project. Negative pressure air conditioning is required for clearance of anaesthetic gases.

## Items in room
Standard Components - Items in room
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
Items | [6 empty] | Service allocations
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment | Electrical | Data | Cold Water | Hot Water | Warm Water | Tap | Drainage | Gas | Steam
BIRM-A | Birthing Room - LDR Without Bath | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | CLGE-001 | CEILING: plasterboard, bulkhead, paint | 1 | above joinery items, joinery bulkhead can alternatively be provided | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 4 | optional | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | DWSC-021 | DOOR: 1400mm c/o, 1 1/2 leaf, solid core, paint | 1 | with latchset and hold open device | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | DWWD-011 | CURTAIN: window, drapes | 1 | set | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | DWWH-005 | WINDOW | 2 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ELBO-015 | SWITCH: light, single | 4 | to ensuite, wall panel, door, store | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | for ceiling mounted display screen (TV) | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ELGP-109 | GPO: single, wall mounted | 4 | patient side | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ELGP-122 | GPO: emergency power, single, on services panel | 2 | to services panel, infant resuscitaire | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 4 | staff side | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ELGP-205 | GPO: double, on services panel | 2 | to services panel, infant resuscitaire | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ELGP-208 | GPO: double, wall mounted | 3 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ELGP-208 | GPO: double, wall mounted | 1 | low level, for electric bed | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ELPR-006 | RCD: residual current device | 4 | 1 to staff side, 1 to patient side, 2 to infant resuscitaire | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1 | to infant resuscitaire | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIDI-014 | DISPENSER: disposable glove | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIHR-004 | BRACKET: for display screen, ceiling mounted | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIHR-016 | BRACKET: for sharps bin | 1 | mounted to trolley | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIHR-019 | BRACKET: for suction bottle | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIJO-162 | CUPBOARD: tall, single door, to conceal services | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIJO-209 | CUPBOARD: tall, single door, adj shelves and clothes rail, recessed | 1 | wardrobe | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIRT-030 | TRACK: curtain, flat pelmet | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FIRT-032 | TRACK: curtain, privacy, L shaped | 1 | to door entry | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQBS-001 | BASSINET | 1 | optional, if used for post-natal care | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1 | to adjustable height desk | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQBS-030 | CHAIR: visitor, patient | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQBS-041 | LOUNGE: 2.5 seat | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQCW-006 | BIN: waste, clinical, 240L | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQCW-043 | BIN: waste, sharps, on trolley | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQDW-039 | DESK: adj height, 800W x 600D x 1000H nom. | 1 | recommended adjustable range 900-1100mm high | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQGE-010 | CURTAIN: privacy screen | 1 | to door to room | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQGE-010 | CURTAIN: privacy screen | 1 | optional, to door to equipment store | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQSM-016 | TROLLEY | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQSN-051 | LOCKER: 2 drawers, open under, lockable, mobile | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQTA-013 | TABLE: overbed | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | FQTA-053 | TABLE: round, 600 diam, low height | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1 | refer to BWHS-A (similar) | No | No | Yes | No | Yes | Yes | Yes | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 |  | No | No | Yes | No | Yes | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | ITBU-012 | DISPLAY SCREEN: patient entertainment system | 1 |  | Yes | Yes | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ITCL-004 | BUTTON: nurse call, emergency | 2 | staff side and at infant resus, + indicator button & light | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | patient side, + indicator button & light | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | staff side, + indicator button & light | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | patient side, with light switch | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | patient side | No | Yes | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | above bench, 1 outlet for telephone | No | Yes | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | ITIN-070 | OUTLET: MATV | 1 | may be replaced with data outlet depending on entertainment system | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 |  | Yes | Yes | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1 |  | No | Yes | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1 |  | No | Yes | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | LIFX-011 | LIGHT: reading light | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | LIFX-017 | LIGHT: night light | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1 | optional | Yes | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | LIPO-006 | LIGHT: examination, mobile | 1 |  | Yes | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | to services panel, infant resuscitaire | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | MGAS-023 | OUTLET: medical air (MA), wall mounted | 1 | in cupboard | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | MGAS-033 | OUTLET: nitrous oxide (N2O), wall mounted | 1 | in cupboard | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | to services panel, infant resuscitaire | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | MGAS-043 | OUTLET: oxygen (O2), wall mounted | 2 | in cupboard | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | MGAS-055 | OUTLET: scavenge, wall mounted | 1 | in cupboard | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | to services panel, infant resuscitaire | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | MGAS-063 | OUTLET: suction, wall mounted | 1 | in cupboard | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | MGFP-002 | FLOWMETER: medical air | 2 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | MGFP-004 | FLOWMETER: oxygen | 3 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | MMBE-005 | BED: patient, birthing, electric | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | MMGE-002 | ADAPTER: suction | 2 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | MMGE-063 | MONITOR: patient, foetal, CTG | 1 | mobile | Yes | Yes | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | MMGE-100 | SUCTION BOTTLE | 2 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | MMMA-011 | MIXER: oxygen and nitrous oxide | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | MMMA-016 | RESUSCITAIRE: infant, wall mounted | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FFE | MMMA-017 | MANTLEPIECE: birthing, wall mounted | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | to infant resuscitaire | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | WLPR-002 | WALL PROTECTION: bed protection wall panel, composite, laminate finish | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | WLPR-011 | WALL PROTECTION: corner guards, to 900 AFFL min. | 5 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1 |  | No | No | No | No | No | No | No | No | No
BIRM-A | Birthing Room - LDR Without Bath | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin | No | No | No | No | No | No | No | No | No