## RDS
Standard Components - Room Data Sheet
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
Briefing
|  |  | Remarks | Occupancy | Hours of Operation | Acoustics | Description / Special Requirements
AusHFG Room Code | Room Name | Prog. Area | (Multi line edit) | (Single line edit) | (Single line edit) | (Single line edit) | (Formatted text)
BPATH | Bay - Pathology | 3 |  | 1 - 2 staff | 24 Hours |  | A space for the preparing and testing a range of specimens. This space has been based on four (4) benchtop POCT devices. This level of provision may be more common in critical care areas. The space is dependent on the number of devices used in each clinical situation. In a multi-device bay, 300mm gap between each device is recommended. Design to consider height clearances if above bench joinery is installed. This space to be colocated with a pneumatic tube bay and in an area that is supervised by staff. Assumes that staff will access a hand hygiene bay nearby. The workspace will suit a staff member who is standing. Design consideration as to the weight of some POCT devices.

## Items in Room
Standard Components - Items in Room
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
[7 empty] | Fabrics |  |  | Service allocations
Connections
[7 empty] | Material | Finish | Specification | Ele | Data | CdW | HtW | Tap | WmW | Dns | Gas | Stm
AusHFG Room Code | Room Name | Item list | BIM ID | Name | Gross quantity | Comment | (Single line edit) | (Single line edit) | (Single line edit) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field)
BPATH | Bay - Pathology | FE | FE-11400 | Computer - PC & monitor | 1 |  |  |  |  | Yes | Yes | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FE | FE-43800 | Refrigerator - under-bench, 120 Litre | 1 | Local temperature monitoring. Optional, depends on frequency of collection. |  |  |  | Yes | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FE | FE-54510 | Telephone - handset, standard | 1 | Bench or wall mounted, optional. |  |  |  | No | Yes | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FE | FE-6459 | Point-of-Care device | 1 | Number to be project specific. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FE | FE-41150 | Printer - labels | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FF | FF-1050 | Bench - laminate | 1 | With optional inset sink. Length to suit layout, refer RLS or project specific requirements. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FF | FF-12900 | Wire panel - for storage-bins | 1 | Over bench. Number to suit layout. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FF | FF-1470 | Bracket - sharps | 1 | Optional. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FF | FF-19700 | Sharps bin | 1 | Optional. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FF | FF-21755 | Storage bins | 2 | Number to suit. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FF | FF-4410 | Cupboard - under bench | 2 | With adjustable shelf. Number to suit layout, refer RLS or project specific requirements. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FF | FF-4470 | Cupboard - wall mounted, lockable (s) | 4 | Over bench. Number to suit layout, refer RLS or project specific requirements. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FF | FF-24124 | Waste bin - clinical, 20 litres | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | FF | FF-5830 | Dispenser - antiseptic hand rub | 1 | May not be required if Hand Hygiene Bay colocated. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | SE | S-9001 | GPO: Double | 4 | Number depends on project specific requirements. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | SE | S-9011 | GPO: Emergency power, double (H) | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | SE | S-1006 | Voice / Data outlet: double | 4 | Number depends on project specific requirements. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | SE | S-5000 | Airconditioning | 1 | Consistent with adjoining space. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | SE | S-6000 | General fluorescent | 1 | Consistent with adjoining space. |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | SE | S-6055 | Special: task light, built in | 2 | under O/H cupboards |  |  |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | Fabrics | 1000 | Floor finish | 1 |  | Vinyl | Standard slip resistant | Seamless, coved | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | Fabrics | 3000 | Wall finish | 1 |  | Paint | Acrylic, washable |  | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | Fabrics | 2670 | Wall protection | 1 | Align heights of corner guards with other corridor elements |  |  | Corner guards to 900 AFFL minimum | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | Fabrics | 5007 | Ceiling - Plasterboard | 1 | Between above bench joinery and ceiling. | Plasterboard | Paint | Bulkhead | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | Fabrics | 5010 | Ceiling - Plasterboard | 1 | Acoustic tiles acceptable | Plasterboard | Paint, washable | Flush set, suspended | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | Fabrics | 6030 | Cornice | 1 |  | Aluminium | Powdercoat | 24mm shadow angle | No | No | No | No | No | No | No | No | No
BPATH | Bay - Pathology | Fabrics | 2000 | Skirting | 1 |  | Vinyl | Prefinished | Floor vinyl coved, 150 mmH | No | No | No | No | No | No | No | No | No