## Room Data
Standard Components - Room Data Sheet
Date of issue: 13.05.2020
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
CLCR-I | Catheter Laboratory Interventional Control Room | 14 | Available 24 Hours | Up to 3 staff |  | Acoustic privacy is required | Room adjacent to Catheter Lab Interventional Procedure Room/s with direct visibility to the patient for remote control of equipment, and review and reporting of images. The Control Room may serve two procedure rooms. Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. Cable ducts from Catheter Lab Procedure Room and Equipment room will be required.

## Items
Standard Components - Items in room
Date of issue: 13.05.2020
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
CLCR-I | Catheter Laboratory Interventional Control Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 6
CLCR-I | Catheter Laboratory Interventional Control Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
CLCR-I | Catheter Laboratory Interventional Control Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
CLCR-I | Catheter Laboratory Interventional Control Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
CLCR-I | Catheter Laboratory Interventional Control Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | to procedure room
CLCR-I | Catheter Laboratory Interventional Control Room | FA | DWGL-002 | DOOR: 910mm c/o, single leaf, half glazed, solid core/glass, paint | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FA | DWGL-006 | DOOR: 910mm c/o, single leaf, observation panel, solid core/lead glass, lead shielding, paint | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 3
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | DWWD-004 | BLIND: roller | 1 | to observation window for patient privacy; refer to local infection control policies
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIJO-143 | CUPBOARD: wall mount, double door, adj shelves | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITBU-013 | DISPLAY SCREEN: patient flow system | 1 | patient journey board
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-009 | COMPUTER: hi performance, for clinical / PACS use | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-023 | PRINTER: networked | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 1 | optional
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 2 | on bench, 1 cordless optional
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITAV-014 | MICROPHONE: surface mounted, gooseneck | 2 | part of communications system integrated with imaging system
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIJO-031 | BENCH: 750D, open under, laminate finish | 2
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIJO-099 | CAC: cable access cap | 5
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIJO-212 | BULKHEAD: joinery | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1 | for reporting station
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | MMMI-369 | DISPLAY SCREEN: contrast injector | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIHR-070 | BRACKET: for display screen, dual, desk mounted | 1
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | FIHR-075 | BRACKET: for computer CPU, underdesk mounted | 3
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | MMMI-386 | CONTROL UNIT: c/w display, control box and keyboard, Cath Lab | 2
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-068 | COMPUTER: hi performance, for clinical / PACS use, CPU, underbench mount | 3
CLCR-I | Catheter Laboratory Interventional Control Room | FFE | ITNE-069 | DISPLAY SCREEN: PACS, portrait | 2
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 6 | 2 outlets for telephone
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELPR-006 | RCD: residual current device | 3
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELBO-001 | BUTTON: emergency stop | 1 | optional
CLCR-I | Catheter Laboratory Interventional Control Room | SE | MEGE-007 | CONTROL PANEL: temperature | 1 | for control room temperature, optional
CLCR-I | Catheter Laboratory Interventional Control Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 1
CLCR-I | Catheter Laboratory Interventional Control Room | SE | MGFP-001 | ALARM PANEL: medical gases | 1
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 8
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
CLCR-I | Catheter Laboratory Interventional Control Room | SE | ELGP-130 | GPO: UPS, single | 7