## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
CTIR | CT Imaging Room | 45 | Available 24 Hours | 1 patient; up to 3 staff assisting |  | Acoustic privacy is required | A Room for Computed Tomography (CT) scanning procedures with an adjacent Control Room and Equipment Room. The Control Room and Equipment room will need to be located in close proximity due to acceptable cable lengths. The Control Room may serve more than one CT Imaging Room. General Anaesthesia may be administered in this room. Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. Unistrut ceiling able to carry moving loads may be required. Cable ducts between Control Room and Equipment Room will be required. 'X-ray in Use' lights are required outside each entry door. Power should be on emergency supply. Dimmable lighting is optional.

## Items
Standard Components - Items in room
Date of issue: 28.05.2019
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
CTIR | CT Imaging Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin
CTIR | CT Imaging Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 8
CTIR | CT Imaging Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 2 | optional
CTIR | CT Imaging Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 3
CTIR | CT Imaging Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 4
CTIR | CT Imaging Room | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 2 | to basin
CTIR | CT Imaging Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | at 3000 mm AFFL; sound absorbing plasterboard recommended
CTIR | CT Imaging Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | safety glass, to control room
CTIR | CT Imaging Room | FA | DWSC-005 | DOOR: 910mm c/o, single leaf, solid core, lead shielding, paint | 1
CTIR | CT Imaging Room | FA | DWSC-034 | DOOR: 1800mm c/o, double leaf, solid core, lead shielding, paint | 1
CTIR | CT Imaging Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
CTIR | CT Imaging Room | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 8
CTIR | CT Imaging Room | FFE | MGFP-002 | FLOWMETER: medical air | 1
CTIR | CT Imaging Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2
CTIR | CT Imaging Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
CTIR | CT Imaging Room | FFE | FIJO-135 | CUPBOARD: underbench, double door, adj shelves | 1
CTIR | CT Imaging Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
CTIR | CT Imaging Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | near door
CTIR | CT Imaging Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
CTIR | CT Imaging Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
CTIR | CT Imaging Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
CTIR | CT Imaging Room | FFE | FIHR-030 | HOOK: equipment, on mounting plate | 1
CTIR | CT Imaging Room | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1
CTIR | CT Imaging Room | FFE | MMGE-131 | TROLLEY: monitor | 1
CTIR | CT Imaging Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2
CTIR | CT Imaging Room | FFE | MMBE-042 | SLIDE: patient | 1
CTIR | CT Imaging Room | FFE | MMGE-142 | TROLLEY: sharps bin | 1
CTIR | CT Imaging Room | FFE | MMGE-002 | ADAPTER: suction | 2
CTIR | CT Imaging Room | FFE | MMGE-100 | SUCTION BOTTLE | 2
CTIR | CT Imaging Room | FFE | FQCW-005 | BIN: waste, clinical, 20L | 1
CTIR | CT Imaging Room | FFE | FQSM-003 | PEDESTAL: mobile, wire basket unit, 840H | 2
CTIR | CT Imaging Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
CTIR | CT Imaging Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1
CTIR | CT Imaging Room | FFE | MMBE-010 | BED: patient, electric with mattress | 1
CTIR | CT Imaging Room | FFE | ITCL-012 | CAMERA: observation | 2
CTIR | CT Imaging Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
CTIR | CT Imaging Room | FFE | MMHA-008 | MONITOR: patient, cardiac with vital signs | 1
CTIR | CT Imaging Room | FFE | MMGE-032 | DISPLAY SCREEN: slave, for patient monitor | 1
CTIR | CT Imaging Room | FFE | FIJO-059 | BENCH TOP: 600D, joinery under, laminate finish | 1
CTIR | CT Imaging Room | FFE | MMMI-099 | DEVICE: imaging, CT | 1
CTIR | CT Imaging Room | FFE | MMMI-105 | TABLE: patient, fixed, CT | 1
CTIR | CT Imaging Room | FFE | MMMI-107 | INJECTOR: contrast, CT, ceiling mounted | 1
CTIR | CT Imaging Room | FFE | FIJO-184 | CUPBOARD: underbench, double door, no shelves, no toeplate, lockable | 1
CTIR | CT Imaging Room | FFE | FIHR-064 | BRACKET: for display screen, arm mounted, rotating | 1
CTIR | CT Imaging Room | FFE | MMSP-035 | PENDANT: single, dual arm, display screen | 1
CTIR | CT Imaging Room | FFE | MMMI-377 | WARMER: contrast | 1
CTIR | CT Imaging Room | FFE | FIHR-078 | BRACKET: for observation camera, wall mounted | 2
CTIR | CT Imaging Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 2
CTIR | CT Imaging Room | SE | ELPR-006 | RCD: residual current device | 2 | on services panel
CTIR | CT Imaging Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel + indicator button & light
CTIR | CT Imaging Room | SE | ELBO-001 | BUTTON: emergency stop | 1
CTIR | CT Imaging Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | on services panel
CTIR | CT Imaging Room | SE | LIFX-002 | LIGHT: illuminated sign, "X-Ray in Use" | 2 | outside room entry, auto on when modality in use
CTIR | CT Imaging Room | SE | ELBO-015 | SWITCH: light, single | 1
CTIR | CT Imaging Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on services panel
CTIR | CT Imaging Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | on services panel
CTIR | CT Imaging Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel
CTIR | CT Imaging Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on services panel
CTIR | CT Imaging Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
CTIR | CT Imaging Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | for clock
CTIR | CT Imaging Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | to walls
CTIR | CT Imaging Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
CTIR | CT Imaging Room | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 3 | for foot pedal recharge
CTIR | CT Imaging Room | SE | ITSE-025 | DOOR RELEASE: access control, remote activation | 1
CTIR | CT Imaging Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
CTIR | CT Imaging Room | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors as specified
CTIR | CT Imaging Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
CTIR | CT Imaging Room | SE | ELGP-133 | GPO: UPS, single, on services panel | 2
CTIR | CT Imaging Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 6 | on services panel
CTIR | CT Imaging Room | SE | ELPR-007 | UPS: status indicator | 1 | on services panel