## RDS
Standard Components - Room Data Sheet
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
Description
|  |  | Briefing |  |  |  | Description or Special Requirements
|  |  | Remarks | Occupancy | Hours of Operation | Acoustics
AusHFG Room Code | Room Name | Prog. Area | (Multi line edit) | (Single line edit) | (Single line edit) | (Single line edit) | (Formatted text)
CTPR | CT Scanning - Procedure Room | 45 |  | 1 patient; up to 3 staff assisting | Available 24 Hours | Acoustic privacy is required | X-ray Room for CT scanning procedures with an adjacent Control Room and Computer Room. The Control Room may serve more than one CT Scanning Room Acoustics: Acoustic privacy is required Body Protected electrical area Additional Design Considerations: - Radiation shielding as advised by Radiation Consultant - Unistrut ceiling able to carry moving loads may be required - Cable ducts between Control Room and Computer Room may be required - &apos;X-ray in Use&apos; lights are required outside each entry door - Power should be on emergency supply

## Items in Room
Standard Components - Items in Room
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
[7 empty] | Fabrics |  |  | Service allocations
Connections
[7 empty] | Material | Finish | Specification | Ele | Data | CdW | HtW | Tap | WmW | Dns | Gas | Stm
AusHFG Room Code | Room Name | Item list | BIM ID | Name | Gross quantity | Comment | (Single line edit) | (Single line edit) | (Single line edit) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field)
CTPR | CT Scanning - Procedure Room | Fabrics | 1000 | Floor finish | 1 |  | Vinyl | Standard slip resistant | Seamless, coved | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | Fabrics | 2000 | Skirting | 1 |  | Vinyl | Prefinished | Floor vinyl coved, 150 mmH | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | Fabrics | 3000 | Wall finish | 1 |  | Paint | Acrylic, washable |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | Fabrics | 3060 | Wall finish | 1 | to basin; tiles also acceptable | Vinyl | Prefinished | Splashback | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | Fabrics | 4020 | Wall protection | 1 | Equal to 'Acrovyn' | Composite | Prefinished PVC | Crash rail at 900 AFFL | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | Fabrics | 5010 | Ceiling - Plasterboard | 1 | at 3000 mm AFFL | Plasterboard | Paint, washable | Flush set, suspended | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | Fabrics | 6030 | Cornice | 1 |  | Aluminium | Powdercoat | 24mm shadow angle | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | Fabrics | 7025 | Observation | 1 | to control room | Lead glass | Clear | Internal window | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | Fabrics | 8010 | Door protection | 1 | Equal to 'Acrovyn' | Composite | Prefinished PVC | Protection plate to 900 AFFL | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | Fabrics | 8040 | Door protection | 1 | Equal to 'Acrovyn' | Composite | Prefinished PVC | To door frame | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | Fabrics | 9130 | Door - 1600mm | 1 | patient entry/ exit 1500mm clear opening | Solid Core | Paint | 1 1/2 leaf, with lead shielding | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FE | FE-3100 | Basin - scrub, Type A | 1 | refer to BHWS-A (sim) |  |  |  | No | No | Yes | No | Yes | Yes | Yes | No | No
CTPR | CT Scanning - Procedure Room | FE | FE-66500 | IMG - CT Scanning unit | 1 |  |  |  |  | Yes | Yes | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FE | FE-68700 | Xray viewer - 2 screen | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FE | FE-9700 | Clock - analogue, electric | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FF | FF-111670 | Cupboard - with glazed doors | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FF | FF-1465 | Bracket - suction bottle | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FF | FF-150 | Air flowmeter | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FF | FF-16250 | Oxygen flowmeter | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FF | FF-21950 | Suction adapter | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FF | FF-22000 | Suction bottle | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FF | FF-5824 | Waste bin - general | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FF | FF-6152 | Dispenser - disposable glove | 1 | near door |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FF | FF-6154 | Dispenser - paper/ hand towel | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FF | FF-6156 | Dispenser - soap | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FF | FF-6158 | Dispenser - soap antiseptic | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-1005 | Voice / Data outlet: single | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-1006 | Voice / Data outlet: double | 2 | 1 outlet on each panel for telephone |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-1040 | Intercom | 1 | handsfree to control room |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-11000 | Wall panel | 3 | patient services |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-12030 | Lead lining | 1 | to walls and doors as specified |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-1951 | Staff assist call | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-5000 | Airconditioning | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-6000 | General fluorescent | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-6110 | Downlights: dimmable | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-6200 | Light switches | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-7000 | Oxygen (O2) | 2 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-7010 | Nitrous Oxide (N2O) | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-7020 | Medical Air (MA) | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-7040 | Suction | 2 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-7070 | Scavenge | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-8010 | Staff/ Nurse assist call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-8080 | Emergency call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-9000 | GPO: Single | 1 | XR Viewer |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-9001650 | GPO:Double RCD | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-9010 | GPO: Emergency power, single (V) | 6 | to patient services panels |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-9011 | GPO: Emergency power, double (H) | 2 | to walls |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-9015 | GPO: Cleaner | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-9040 | Body protected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | S-9060 | GPO: Special: for equipment | 1 | CT equipment |  |  |  | No | No | No | No | No | No | No | No | No