## RDS
Standard Components - Room Data Sheet
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
CTPR | CT Scanning - Procedure Room | 45 | Available 24 Hours | 1 patient; up to 3 staff assisting |  | Acoustic privacy is required | X-ray Room for CT scanning procedures with an adjacent Control Room and Computer Room. The Control Room may serve more than one CT Scanning Room Additional Design Considerations: Radiation shielding as advised by Radiation Consultant. Unistrut ceiling able to carry moving loads may be required. Cable ducts between Control Room and Computer Room may be required. 'X-ray in Use' lights are required outside each entry door. Power should be on emergency supply. Dimmable downlights are optional.

## Items in room
Standard Components - Items in room
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
Items | [6 empty] | Service allocations
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment | Electrical | Data | Cold Water | Hot Water | Warm Water | Tap | Drainage | Gas | Steam
CTPR | CT Scanning - Procedure Room | FA | CLCN-002 | CORNICE: aluminium, 24mm shadow angle, powdercoat | 1 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | at 3000 mm AFFL | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | DWGL-006 | DOOR: 910mm c/o, single leaf, observation panel, solid core/lead glass, lead shielding, paint | 1 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | to control room | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1 | equal to Acrovyn | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | equal to Acrovyn | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | DWSC-025 | DOOR: 1500mm c/o, 1 1/2 leaf, solid core, lead shielding, paint | 1 | patient entry/ exit | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ELBO-015 | SWITCH: light, single | 2 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ELGP-106 | GPO: single, special, for equipment | 1 | CT equipment | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | XR Viewer | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 4 | on Services Panel 2 | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 2 | on Services Panel 1 | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ELGP-206 | GPO: double, on services panel, RCD protected | 1 | on Services Panel 1 | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ELGP-206 | GPO: double, on services panel, RCD protected | 1 | on Services Panel 2 | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 2 | to walls | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | near door | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | FIJO-121 | CUPBOARD: tall, double door, sliding and glazed, adj shelves | 1 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | FLVY-009 | FLOOR FINISH: vinyl, seamless, coved, wet area, non-slip | 1 | to basin | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | HYBA-009 | BASIN: handwash, Type A, scrub | 1 | refer BHWS-A (similar) | No | No | Yes | No | Yes | Yes | Yes | No | No
CTPR | CT Scanning - Procedure Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin | No | No | Yes | No | Yes | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on Services Panel 1 + indicator button & light | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on Services Panel 1 + indicator button & light | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ITCL-009 | BUTTON: nurse call, staff / nurse assist | 1 | on Services Panel 1 + indicator button & light | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1 |  | Yes | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1 | on Services Panel 2 | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | 1 outlet on each panel for telephone | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 | handsfree to control room | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on Services Panel 3 | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1 | on Services Panel 3 | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | on Services Panel 3 | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1 | on Services Panel 3 | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | on Services Panel 3 | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | MGFP-002 | FLOWMETER: medical air | 1 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | MMGE-001 | ADAPTER: scavenge | 1 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | MMGE-002 | ADAPTER: suction | 2 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | MMGE-100 | SUCTION BOTTLE | 2 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | MMGE-152 | VIEWER: X-ray, 2 screen | 1 |  | Yes | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FFE | MMMI-103 | SYSTEM: imaging, CT | 1 |  | Yes | Yes | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 3 | patient services | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | WLPR-021 | WALL PROTECTION: crash rail, at 900 AFFL | 1 | Equal to Acrovyn | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors as specified | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1 |  | No | No | No | No | No | No | No | No | No
CTPR | CT Scanning - Procedure Room | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin; tiles also acceptable | No | No | No | No | No | No | No | No | No