## RDS
Standard Components - Room Data Sheet
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
Description
|  |  | Briefing |  |  |  | Description or Special Requirements
|  |  | Remarks | Occupancy | Hours of Operation | Acoustics
AusHFG Room Code | Room Name | Prog. Area | (Multi line edit) | (Single line edit) | (Single line edit) | (Single line edit) | (Formatted text)
GENXR | General X-Ray | 30 |  | 1 patient; 1 - 2 staff | Available 24 Hours |  | General X-ray is a room for undertaking general radiographic procedures, which may be located within the Imaging Unit or within an emergency care area. Additional Design Requirements: - Radiation shielding to walls and glazing as advised by Radiation Consultant - Unistrut type ceiling may be required, depending on equipment to be installed - Cable ducting between the X-ray table, vertical bucky and control area may be required, depending on equipment type and layout - &apos;X-ray in Use&apos; warning lights will be required outside each entry door Body Protected electrical area

## Items in Room
Standard Components - Items in Room
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
[7 empty] | Fabrics |  |  | Service allocations
Connections
[7 empty] | Material | Finish | Specification | Ele | Data | CdW | HtW | Tap | WmW | Dns | Gas | Stm
AusHFG Room Code | Room Name | Item list | BIM ID | Name | Gross quantity | Comment | (Single line edit) | (Single line edit) | (Single line edit) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field)
GENXR | General X-Ray | Fabrics | 1000 | Floor finish | 1 |  | Vinyl | Standard slip resistant | Seamless, coved | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | Fabrics | 2000 | Skirting | 1 |  | Vinyl | Prefinished | Floor vinyl coved, 150 mmH | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | Fabrics | 3000 | Wall finish | 1 |  | Paint | Acrylic, washable |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | Fabrics | 3060 | Wall finish | 1 | to basin, tiles also acceptable | Vinyl | Prefinished | Splashback | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | Fabrics | 4030 | Wall protection | 1 | Equal to 'Acrovyn', as requried | Composite | Prefinished PVC | Corner guards to 900 AFFL | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | Fabrics | 5030 | Ceiling - Plasterboard | 1 | at 3000 mm AFFL | Plasterboard | Prefinished | Drop-in tiles, 600x1200 | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | Fabrics | 6030 | Cornice | 1 |  | Aluminium | Powdercoat | 24mm shadow angle | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | Fabrics | 8010 | Door protection | 1 | Equal to 'Acrovyn' | Composite | Prefinished PVC | Protection plate to 900 AFFL | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | Fabrics | 8040 | Door protection | 1 | Equal to 'Acrovyn' | Composite | Prefinished PVC | To door frame | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | Fabrics | 9230 | Doors | 1 |  | Solid Core | Paint | Double leaf, with lead shielding | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FE | FE-2800 | Basin - handwash, Type B | 1 | refer to BHWS-B (sim) |  |  |  | No | No | Yes | No | Yes | Yes | Yes | No | No
GENXR | General X-Ray | FE | FE-54520 | Telephone - handset, handsfree (WM) | 1 |  |  |  |  | No | Yes | No | No | No | No | No | No | No
GENXR | General X-Ray | FE | FE-61600 | IMG - bucky, chest | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FE | FE-62152 | IMG - console, control (integrated stand) | 1 |  |  |  |  | Yes | Yes | No | No | No | No | No | No | No
GENXR | General X-Ray | FE | FE-62400 | IMG - control booth | 1 |  |  |  |  | Yes | Yes | No | No | No | No | No | No | No
GENXR | General X-Ray | FE | FE-66000 | IMG - X-ray table, floating top | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FE | FE-66600 | IMG - X-ray unit, general purpose | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FE | FE-68400 | IMG: generator | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FE | FE-68600 | Xray viewer - 1 screen | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-1050 | Bench - laminate | 1 | cupboards under |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-11905 | Rack - lead apron | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-1465 | Bracket - suction bottle | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-150 | Air flowmeter | 1 | to services panel |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-16250 | Oxygen flowmeter | 1 | to services panel |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-1650 | Chair - client | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | FF-21950 | Suction adapter | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-22000 | Suction bottle | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-4410 | Cupboard - under bench | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-5824 | Waste bin - general | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-6152 | Dispenser - disposable glove | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-6154 | Dispenser - paper/ hand towel | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-6156 | Dispenser - soap | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | FF | FF-6158 | Dispenser - soap antiseptic | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-1006 | Voice / Data outlet: double | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-1040 | Intercom | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-11000 | Wall panel | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-12030 | Lead lining | 1 | to walls and doors as specified |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-1951 | Staff assist call | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-5000 | Airconditioning | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-6001 | General: colour corrected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-6110 | Downlights: dimmable | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-6200 | Light switches | 3 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-7000 | Oxygen (O2) | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-7020 | Medical Air (MA) | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-7040 | Suction | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-8010 | Staff/ Nurse assist call | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-8080 | Emergency call | 1 | patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-9000 | GPO: Single | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-9001 | GPO: Double | 4 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-9010 | GPO: Emergency power, single (V) | 2 | to patient services panel |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-9010 | GPO: Emergency power, single (V) | 4 | 2-patient services panel;2-walls |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-9040 | Body protected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
GENXR | General X-Ray | SE | S-9060 | GPO: Special: for equipment | 2 | Xray Equipment, chest bucky |  |  |  | No | No | No | No | No | No | No | No | No