## Room Data
Standard Components - Room Data Sheet
Date of issue: 28.06.2021
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
LINAC-CR | Linear Accelerator Control Room | 21 | 8am to 6pm, Monday – Friday (extended hours, weekends and/or emergency access may be required) | 2-5 staff |  | Acoustic privacy is required | The Linear Accelerator Control room is located adjacent to the Treatment Room for remote control of equipment with audio and visual contact with the patient via cameras, speakers and microphones. Each Treatment Room requires a Control Room. Cable pathways for services between the Control room to the Treatment Room must be established and coordinated prior to construction. Radiation shielding is to be as advised by a Radiation Consultant. Additional Considerations: Air conditioning to provide comfort conditions; high equipment loads shall be accounted for. The need for and extent of emergency power and UPS must be reviewed to suit site specific requirements. Small UPS units are normally required as a minimum to support control and operating equipment. In some cases, these are supplied with the medical equipment. Other emergency power and UPS requirements to be confirmed based on risk assessment considering the impact of a power outage on patient care/safety. Cable management and the potential for computers to be consolidated into a rack (either within the room or in a dedicated, collocated air conditioned room/cupboard) to be determined at project level. Provision of a height adjustable workstation may be considered depending on local policies and staff accessibility requirements.

## Items
Standard Components - Items in room
Date of issue: 28.06.2021
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
LINAC-CR | Linear Accelerator Control Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1
LINAC-CR | Linear Accelerator Control Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 9
LINAC-CR | Linear Accelerator Control Room | FA | DWPR-001 | DOOR PROTECTION: kickplate, to 300 AFFL | 1 | optional
LINAC-CR | Linear Accelerator Control Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 5
LINAC-CR | Linear Accelerator Control Room | FA | CLTI-002 | CEILING: acoustic, drop-in tiles, 600 x 1200 | 1
LINAC-CR | Linear Accelerator Control Room | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | optional
LINAC-CR | Linear Accelerator Control Room | FA | CLCN-009 | CORNICE: wall trim fixing | 1
LINAC-CR | Linear Accelerator Control Room | FFE | FQBS-018 | CHAIR: office, ergonomic | 5
LINAC-CR | Linear Accelerator Control Room | FFE | FQGE-005 | CLOCK: analogue, battery operated | 1
LINAC-CR | Linear Accelerator Control Room | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 5
LINAC-CR | Linear Accelerator Control Room | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1
LINAC-CR | Linear Accelerator Control Room | FFE | ITAV-057 | SPEAKER: ceiling recessed | 1 | may be bench mounted, or integrated with microphone
LINAC-CR | Linear Accelerator Control Room | FFE | ITNE-049 | TELEPHONE: handset, desktop | 1
LINAC-CR | Linear Accelerator Control Room | FFE | ITAV-014 | MICROPHONE: surface mounted, gooseneck | 1
LINAC-CR | Linear Accelerator Control Room | FFE | FIJO-031 | BENCH: 750D, open under, laminate finish | 2
LINAC-CR | Linear Accelerator Control Room | FFE | FIJO-099 | CAC: cable access cap | 8
LINAC-CR | Linear Accelerator Control Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead is also acceptable
LINAC-CR | Linear Accelerator Control Room | FFE | FIJO-213 | PANEL: joinery carcass end | 2
LINAC-CR | Linear Accelerator Control Room | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 2 | to patient observation monitors
LINAC-CR | Linear Accelerator Control Room | FFE | ITNE-067 | KEYBOARD: keyboard and mouse only | 1 | for patient observation cameras to control pan, tilt and zoom; may be custom control depending on camera system
LINAC-CR | Linear Accelerator Control Room | FFE | FIHR-080 | BRACKET: for bench, 750D, wall mounted | 5 | legs are also acceptable, quantity to be confirmed to achieve required support, ensure location does not impact leg clearance at workstations
LINAC-CR | Linear Accelerator Control Room | FFE | ITNE-071 | COMPUTER: desktop, with dual display screens | 2 | for access to local network including patient records and images for planning and reporting
LINAC-CR | Linear Accelerator Control Room | FFE | MMRA-006 | ACCESS PANEL: medical physics port | 1
LINAC-CR | Linear Accelerator Control Room | FFE | MMRA-009 | COMPUTER: linear accelerator control system PC | 6 | quantity to be confirmed based on vendor and equipment requirements; current provision allows for 1 x linear accelerator control system, 1 x record and verify system, and 2-4 x imaging and patient set up systems (e.g. gating and patient positioning).
LINAC-CR | Linear Accelerator Control Room | FFE | MMRA-010 | CONSOLE: control, linear accelerator | 1
LINAC-CR | Linear Accelerator Control Room | FFE | MMRA-007 | CABINET: linear accelerator computer cabinet | 1 | size, configuration and location to be confirmed by vendor
LINAC-CR | Linear Accelerator Control Room | FFE | ITNE-074 | DISPLAY SCREEN: patient observation | 2
LINAC-CR | Linear Accelerator Control Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3
LINAC-CR | Linear Accelerator Control Room | SE | ITIN-036 | OUTLET: data, triple RJ45, wall mounted | 8 | quantity to suit computer and display screen requirements
LINAC-CR | Linear Accelerator Control Room | SE | ELPR-006 | RCD: residual current device | 3
LINAC-CR | Linear Accelerator Control Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1
LINAC-CR | Linear Accelerator Control Room | SE | ELBO-001 | BUTTON: emergency stop | 1 | may be integral to linear accelerator control console
LINAC-CR | Linear Accelerator Control Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1
LINAC-CR | Linear Accelerator Control Room | SE | ELGP-208 | GPO: double, wall mounted | 10
LINAC-CR | Linear Accelerator Control Room | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 14 | quantity to suit computer and display screen requirements
LINAC-CR | Linear Accelerator Control Room | SE | ELGP-106 | GPO: single, special, for equipment | 1 | to linear accelerator computer cabinet
LINAC-CR | Linear Accelerator Control Room | SE | ITSE-006 | BUTTON: duress, fixed | 1
LINAC-CR | Linear Accelerator Control Room | SE | ELBO-012 | SWITCH: light, dimmer | 1
LINAC-CR | Linear Accelerator Control Room | SE | ITIN-055 | OUTLET: data, USB, wall mounted | 1 | connection for music to bunker