## Room Data
Standard Components - Room Data Sheet
Date of issue: 13.05.2020
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | 16.5 | 24 Hours | 1 baby; 4-5 staff; 1-2 carers |  | Sound attenuation level high | A single bay for neonates requiring intensive/ high dependency nursing care and medical treatment, that may include monitoring, invasive or non-invasive ventilation (CPAP), nitric treatment, surgical intervention or cooling therapy. The bay will be in close proximity to a Type A scrub basin, shared between bays. The bay will require access for a mother in a bed and a transport trolley. Facilities for storing parent/carer's belongings is required. An external outlook is desirable from a group of neonatal bays. Lighting will require special consideration and should be dimmable, zoned in parent spaces, and direct overhead lighting should be avoided.

## Items
Standard Components - Items in room
Date of issue: 13.05.2020
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | low VOC vinyl, consider acoustic options
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 5
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FA | WLPR-006 | WALL PROTECTION: corner guards | 2
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 3
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 | consider acoustic properties
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FA | WLWA-025 | WALL FINISH: vinyl, to 1200 AFFL | 3
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MGFP-002 | FLOWMETER: medical air | 2 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FIHR-019 | BRACKET: for suction bottle | 4
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FIHR-016 | BRACKET: for sharps bin | 1 | mounted on trolley FQSM-035
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FQBS-019 | CHAIR: office, ergonomic, high | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FQBS-028 | CHAIR: visitor, clinical areas | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 2 | shared
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FQCW-043 | BIN: waste, sharps, on trolley | 1 | mounted on trolley FQSM-035
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FIJO-194 | SHELF: laminate | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MMGE-002 | ADAPTER: suction | 4 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MMGE-100 | SUCTION BOTTLE | 4
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FQSM-035 | TROLLEY: small | 2
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FIBM-018 | WHITEBOARD: fixed | 1 | may be optionally replaced with a care board
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FQCA-086 | WARMER: bottle, mobile | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MMGE-076 | PUMP: breast | 1 | on trolley/stand
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | LIFX-013 | LIGHT: procedure, ceiling mounted, small | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | on trolley
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MMMA-007 | INCUBATOR: infant, humidicrib | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MMHA-005 | MONITOR: patient, cardiac | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MMGE-065 | MONITOR: patient, neonatal | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MMSP-018 | PENDANT: single, wet and dry combination (fixed or adjustable height) | 1 | optional; would replace wall mounted service panel MMSP-026
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | ITNE-050 | TELEPHONE: handset, desktop, handsfree | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MMHA-024 | VENTILATOR: continuous positive airway pressure (CPAP) | 1 | optional; provide bracket
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MMHA-025 | VENTILATOR: neonatal | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FQBS-055 | CHAIR: overnight fold out | 1 | compatible with kangaroo care, breastfeeding
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FIHR-067 | BRACKET: for monitor, wall mounted, articulated, adj height | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FQSM-044 | CUPBOARD: mobile, double door, adjustable shelves | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FQGE-091 | FRIDGE: milk, small | 1 | optional
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | ITCL-036 | CAMERA: neonatal | 1 | optional; must be able to be turned off within the bay; storage and use of images to be considered
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | MMMA-018 | MIXER: oxygen and air | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FIHR-083 | BRACKET: for neonatal camera, crib mounted | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FIHR-084 | BRACKET: for oxygen and air mixer | 1 | optional
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FIHR-085 | BRACKET: for neonatal monitor, crib mounted | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | FFE | FIHR-086 | BRACKET: for CPAP ventilator, wall mounted | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELPR-006 | RCD: residual current device | 2 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELGP-101 | GPO: single, ceiling mounted | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | LIFX-017 | LIGHT: night light | 1 | switch to be located at cubicle or staff station
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 4 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 4 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | MGAS-062 | OUTLET: suction, on services panel | 4 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | on wall; for parent/carer use
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | on wall; for parent/carer use
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELGP-208 | GPO: double, wall mounted | 1 | to parent/carer area; may include optional USB charging point
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELGP-122.1 | GPO: emergency power, single, on services panel (V) | 12 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELGP-223 | GPO: emergency power, double, wall mounted | 4
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELPR-004 | MONITOR: line isolation | 1 | 1 required for UPS
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 2
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | MMSP-026 | SERVICES PANEL: wall mounted | 3 | flush mounted
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELBO-005 | EQUIPOTENTIAL JUNCTION BOX | 1
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELBO-012 | SWITCH: light, dimmer | 2
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ITIN-035 | OUTLET: data, triple RJ45, on services panel | 2 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELGP-133 | GPO: UPS, single, on services panel | 4 | on services panel
NBIC-HD | Neonatal Bay - Intensive Care/ High Dependency Care | SE | ELPR-007 | UPS: status indicator | 1 | on services panel