## RDS
Standard Components - Room Data Sheet
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
Description
|  |  | Briefing |  |  |  | Description or Special Requirements
|  |  | Remarks | Occupancy | Hours of Operation | Acoustics
AusHFG Room Code | Room Name | Prog. Area | (Multi line edit) | (Single line edit) | (Single line edit) | (Single line edit) | (Formatted text)
ORGN | Operating Room - General | 42 |  | 1 patient; 5 - 8 staff | 24 Hours |  | The Operating Room - General provides an aseptic environment in which to carry out surgical procedures under local, regional or general anaesthetic. The General Operating Room may be used for General surgery and specialties including ENT, Urology, Gynaecology, Ophthalmology, Plastic Surgery and any other procedures that do not require bulky equipment. The Operating Room will have direct access to a Scrub-up Room, exit bay and sterile stock/set-up area. Cardiac Protected electrical area Special Air Conditioning - HEPA filtration Additional Design Considerations: - Minimum critical dimension 6000mm. - &apos;Laser In Use&apos; warning lights may be required outside all exit doors to room if laser equipment is to be used - Lead shielding may be required depending on use of X-ray equipment - to be assessed by Radiation Consultant - External windows are desirable - double glazed with integral Venetian blinds

## Items in Room
Standard Components - Items in Room
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
[7 empty] | Fabrics |  |  | Service allocations
Connections
[7 empty] | Material | Finish | Specification | Ele | Data | CdW | HtW | Tap | WmW | Dns | Gas | Stm
AusHFG Room Code | Room Name | Item list | BIM ID | Name | Gross quantity | Comment | (Single line edit) | (Single line edit) | (Single line edit) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field)
ORGN | Operating Room - General | Fabrics | 1000 | Floor finish | 1 |  | Vinyl | Standard slip resistant | Seamless, coved | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 2000 | Skirting | 1 |  | Vinyl | Prefinished | Floor vinyl coved, 150 mmH | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 3010 | Wall finish | 1 |  | Paint | Scrub resistant | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 3038 | Wall finish | 1 |  | Vinyl | Prefinished | Wall vinyl to 900 AFFL | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 5010 | Ceiling - Plasterboard | 1 |  | Plasterboard | Paint, washable | Flush set, suspended | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 6015 | Cornice | 1 |  | Plasterboard | Paint | Flush coved between wall & ceiling | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 7010 | Observation | 1 | with cover (sliding) | Glass | Clear | Through door/s | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 7035 | Observation | 1 | double glazing with integral venetian, optional | Alum/ Glass | Powdercoat | Internal window | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 8010 | Door protection | 1 | Equal to 'Acrovyn' | Composite | Prefinished PVC | Protection plate to 900 AFFL | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 8040 | Door protection | 1 | Equal to 'Acrovyn' | Composite | Prefinished PVC | To door frame | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 9004 | Door - 1100mm | 1 | 910mm clear opening to set-up/scrub rooms | Solid Core/ Glass | Paint | Single leaf, observation panel | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 9004 | Door - 1100mm | 1 | 1100mm clear opening, to clean-up room | Solid Core/ Glass | Paint | Single leaf, observation panel | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | Fabrics | 9225 | Door - 1500mm | 1 | 1500mm clear opening to patient entry/exit | Solid Core/ Glass | Paint | Double leaf, observation panel | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FE | FE-11400 | Computer - PC & monitor | 1 | optional |  |  |  | Yes | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-16000 | Diathermy machine | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-2795 | Television (large flatscreen) | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-32600 | Light - operating, 2 arm, 950/550 diam nominal | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-35960 | Monitor - anaesthetic | 1 | with printer, on anaesth machine |  |  |  | Yes | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-47000 | Services pendant - ceiling | 2 | cable support arm optional |  |  |  | Yes | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-54520 | Telephone - handset, handsfree | 1 | to bench |  |  |  | No | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-55750 | Tourniquet - surgical | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FE | FE-68900 | Xray viewer - 4 screen | 1 | recessed, 4x1 layout |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-700 | Anaesthetic machine | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-9750 | Clock - slave, analogue | 1 | recessed |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FE | FE-9800 | Clock - time elapse | 1 | beside analogue, recessed |  |  |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General | FF | FF-10200 | Trolley - instrument | 4 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-1102 | Bench - resinate, mobile | 1 | for computer, optional | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-11100 | IV stand - mobile | 2 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-11500 | Kickbucket & stand | 2 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-12300 | Linen carrier - dirty, single | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-1252 | Blind - Venetian integral | 1 | to internal window, if provided | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-13500 | Table - mayo | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-1440 | Bowl stand - single | 2 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-1465 | Bracket - suction bottle | 4 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-150 | Air flowmeter | 2 | to services pendants | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-15100 | Bracket - monitor | 1 | optional, for TV monitor if required | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-1580 | Chair - anaesthetist | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-16000 | Operating table - manual | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-16250 | Oxygen flowmeter | 4 | to services pendants | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-21730 | Stool - operating, mobile | 2 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-21950 | Suction adapter | 4 | to services pendants | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-22000 | Suction bottle | 4 | in mobile carrier | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-22050 | Suction bottle carrier | 1 | mobile | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-24405 | Whiteboard - magnetic | 1 | 900x900, optional | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-300 | Trolley - anaesthetic | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-6152 | Dispenser - disposable glove | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | FF | FF-8390 | Trolley - gowns | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-1006 | Voice / Data outlet: double | 1 | to bench; 1 outlet for telephone | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-1006 | Voice / Data outlet: double | 2 | 1 / each pendants | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-11000 | Wall panel | 4 | flush mounted | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-1120 | Monitor connection facilities and cabling | 1 | to wall TV monitor, optional | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-1416 | Light Control Panel | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-1951 | Staff assist call | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-5005 | Air conditioning: special | 1 | HEPA filtered | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-5040 | Pressure: positive | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-6001 | General: colour corrected | 1 | variable lighting level, dimmable | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-6074 | Special: 'Laser In Use' lights | 4 | outside room entries, if laser to be used | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-7000 | Oxygen (O2) | 4 | 2 / each pendant | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-7010 | Nitrous Oxide (N2O) | 2 | 1 / each pendant | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-7020 | Medical Air (MA) | 2 | 1 / each pendant | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-7030 | Tool Air | 2 | 1 / each pendant, controls on wall panel | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-7040 | Suction | 4 | 2 / each pendant | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-7050 | Gas Alarm Panel | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-7060 | Tourniquet air | 1 | and controls, on wall, optional | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-7070 | Scavenge | 2 | 1 / each pendant | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-8010 | Staff/ Nurse assist call | 1 | + indicator button & light | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-8080 | Emergency call | 1 | + indicator button & light | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-8090 | Orderly/ Porter call | 1 | optional | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-9000 | GPO: Single | 4 | on walls | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-9000 | GPO: Single | 8 | 4 per pendant | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-9001 | GPO: Double | 1 | to bench, optional | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-9010 | GPO: Emergency power, single (V) | 1 | to wall TV monitor, optional | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-9010 | GPO: Emergency power, single (V) | 2 | on walls | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-9010 | GPO: Emergency power, single (V) | 8 | 4 per pendant | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-9015 | GPO: Cleaner | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-9030 | Cardiac protected | 1 | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-9050 | Direct connection | 3 | XRV, clocks, lights | bar: Connections Ele to Connections Stm
ORGN | Operating Room - General | SE | S-9060 | GPO: Special: for equipment | 1 | 20 amp, for laser, optional | bar: Connections Ele to Connections Stm