## RDS
Standard Components - Room Data Sheet
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
ORGN | Operating Room - General, 42m2 | 42 | 24 Hours | 1 patient; 5 - 8 staff |  |  | The Operating Room - General provides an aseptic environment in which to carry out surgical procedures under local, regional or general anaesthetic. The General Operating Room may be used for General surgery and specialties including ENT, Urology, Gynaecology, Ophthalmology, Plastic Surgery and any other procedures that do not require bulky equipment. The Operating Room will have direct access to a Scrub-up Room, exit bay and sterile stock/set-up area. Additional Design Considerations: Minimum critical dimension 6000mm. 'Laser In Use' warning lights may be required outside all exit doors to room if laser equipment is to be used. Lead shielding may be required depending on use of X-ray equipment - to be assessed by Radiation Consultant. External windows are desirable - double glazed with integral venetian blinds.

## Items in room
Standard Components - Items in room
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
Items | [6 empty] | Service allocations
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment | Electrical | Data | Cold Water | Hot Water | Warm Water | Tap | Drainage | Gas | Steam
ORGN | Operating Room - General, 42m2 | FA | CLCN-007 | CORNICE: plasterboard, flush coved between wall and ceiling, paint | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | DWGL-005 | DOOR: 910mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | 910mm clear opening to set-up/scrub rooms, quantity to suit layout | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | DWGL-020 | DOOR: 1100mm c/o, single leaf, observation panel, solid core/glass, paint | 1 | 1100mm clear opening, to clean-up room, quantity to suit layout | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | DWGL-042 | DOOR: 1500mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1 | 1500mm clear opening to patient entry/exit, quantity to suit layout | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | DWOB-003 | OBSERVATION: internal window, alum/glass, double glazed, integral venetian blind, powdercoat | 1 | optional | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | DWOB-010 | OBSERVATION: through door/s, glass, clear | 1 | with cover (sliding) | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1 | Equal to Acrovyn | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | Equal to Acrovyn | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | DWWD-007 | BLIND: venetian, integral | 1 | to internal window, if provided | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELBO-003 | CONTROL PANEL: lighting | 1 | to services panel P2 | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELBO-005 | EQUIPOTENTIAL JUNCTION BOX | 1 | to services panel P1 | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELBO-009 | SWITCH: "Laser in use" | 1 | to services panel P4 | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELGE-001 | DIRECT CONNECTION: power | 3 | XRV, clocks, lights | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELGP-104 | GPO: single, on pendant | 8 | 4 per pendant | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELGP-105 | GPO: single, on services panel | 2 | to services panel P2 | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELGP-105 | GPO: single, on services panel | 2 | to services panel P3, 1 no. 20 amp for laser (optional) | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELGP-121 | GPO: emergency power, single, on pendant | 8 | 4 per pendant | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELGP-122 | GPO: emergency power, single, on services panel | 3 | to services panel P1 | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELGP-123 | GPO: emergency power, single, wall mounted | 1 | to wall TV display screen. optional | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELGP-209 | GPO: double, wall mounted, RCD protected | 1 | to bench, optional | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ELPR-006 | RCD: residual current device | 3 | to services panel P1, P2 and P3 | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | FIBM-019 | WHITEBOARD: fixed, magnetic | 1 | 900x900, optional | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | FIDI-014 | DISPENSER: disposable glove | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | FIHR-006 | BRACKET: for display screen, wall mounted | 1 | optional, for TV display screen if required | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | FIHR-019 | BRACKET: for suction bottle | 4 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | FIJO-022 | BENCH: 600D, open under, resinate finish, mobile | 1 | for computer, optional | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | ITBU-011 | DISPLAY SCREEN: MATV | 1 | for television, could be replaced with IPTV, to be decided at project level | Yes | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | to services panel P4, + indicator button & light | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ITCL-005 | BUTTON: nurse call, orderly / porter | 1 | to services panel P4, optional | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | to services panel P4, + indicator button & light | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | ITCL-016 | CLOCK: clock system slave, analogue | 1 | recessed | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | ITCL-022 | CLOCK: clock system slave, elapsed time | 1 | beside analogue, recessed | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ITCL-023 | CONTROL PANEL: master clock system | 1 | to services panel P2 | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to wall TV display screen, optional, with monitor connection facilities and cabling | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ITIN-024 | OUTLET: data, double RJ45, on pendant | 2 | 1 to each pendant | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1 | to bench; 1 outlet for telephone | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | ITNE-007 | COMPUTER: desktop, with display screen | 1 | optional | Yes | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | ITNE-050 | TELEPHONE: handset, desktop, handsfree | 1 | to bench | No | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | LIFX-003 | LIGHT: illuminated sign, "Laser in Use" | 4 | outside room entries, if laser to be used | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | MESU-013 | SUPPLY AIR: grille, HEPA filtered | 4 | low level | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | MGAS-021 | OUTLET: medical air (MA), on pendant | 2 | 1 on each pendant | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | MGAS-031 | OUTLET: nitrous oxide (N2O), on pendant | 2 | 1 on each pendant | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | MGAS-041 | OUTLET: oxygen (O2), on pendant | 4 | 2 on each pendant | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | MGAS-052 | OUTLET: scavenge, on pendant, with adapter | 2 | 1 on each pendant | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | MGAS-061 | OUTLET: suction, on pendant | 4 | 2 on each pendant | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | MGAS-074 | OUTLET: tool air, on pendant | 2 | 1 on each pendant, controls on wall mounted services panel | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | MGAS-080 | OUTLET: tourniquet air, with controls | 1 | to services panel P1, and controls, optional | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | MGFP-001 | ALARM PANEL: medical gases | 1 | to services panel P1 | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MGFP-002 | FLOWMETER: medical air | 2 | to services pendants | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MGFP-004 | FLOWMETER: oxygen | 4 | to services pendants | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | MGFP-007 | CONTROL PANEL: tool air | 1 | to services panel P1 | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMGE-002 | ADAPTER: suction | 4 | to services pendants | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMGE-014 | CHAIR: anaesthetist | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMGE-045 | IV STAND: mobile | 2 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMGE-048 | LINEN CARRIER: dirty, single | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMGE-100 | SUCTION BOTTLE | 4 | in mobile carrier | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMGE-122 | TROLLEY: gowns | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMGE-123 | TROLLEY: instrument | 4 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMGE-156 | VIEWER: X-ray, 4 screen | 1 | recessed, 4x1 layout | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMHA-001 | CARRIER: suction bottle | 1 | mobile | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMHA-004 | MONITOR: patient, anaesthetic | 1 | with printer, on anaesth machine | Yes | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMOP-028 | KICKBUCKET & STAND: | 2 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMOP-040 | MACHINE: anaesthetic | 1 |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMOP-043 | MACHINE: electrosurgical, diathermy | 1 |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMOP-069 | STAND: bowl, single | 2 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMOP-071 | STOOL: operating, mobile | 2 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMOP-076 | TABLE: mayo | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMOP-079 | TABLE: operating, manual | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMOP-084 | TOURNIQUET: surgical | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMOP-085 | TROLLEY: anaesthetic | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMSP-015 | PENDANT: single, anaesthetic (fixed or adjustable height) | 1 | cable support arm optional | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMSP-016 | PENDANT: single, equipment (fixed or adjustable height) | 1 | surgical, cable support arm optional | Yes | Yes | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | SE | MMSP-026 | SERVICES PANEL: wall mounted | 4 | flush mounted | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FFE | MMSP-028 | LIGHT: operating, 2 arm, 950/550 diam nom. | 1 |  | Yes | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | WLWA-005 | WALL FINISH: paint, scrub resistant | 1 |  | No | No | No | No | No | No | No | No | No
ORGN | Operating Room - General, 42m2 | FA | WLWA-022 | WALL FINISH: vinyl, to 900 AFFL | 1 |  | No | No | No | No | No | No | No | No | No