## RDS
Standard Components - Room Data Sheet
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
Description
|  |  | Briefing |  |  |  | Description or Special Requirements
|  |  | Remarks | Occupancy | Hours of Operation | Acoustics
AusHFG Room Code | Room Name | Prog. Area | (Multi line edit) | (Single line edit) | (Single line edit) | (Single line edit) | (Formatted text)
PBTR-R | Patient Bay - Resuscitation | 25 |  | 1 patient; 4 - 5 staff | 24 Hours |  | A Patient Bed Bay for assessment, resuscitation and treatment of patients with critical conditions, located within the Emergency Unit. The Bay may have partition side walls and a bench used for preparation and write-up. Resuscitation and diagnostic equipment should be located in close proximity which may include ventilator, cardiac pacemaking equipment and mobile X-ray machine. Cardiac Protected electrical area

## Items in Room
Standard Components - Items in Room
1st March 2016
Generated by dRofus (C) dRofus Pty Ltd
Items | [6 empty] | Fabrics |  |  | Service allocations
Connections
[7 empty] | Material | Finish | Specification | Ele | Data | CdW | HtW | Tap | WmW | Dns | Gas | Stm
AusHFG Room Code | Room Name | Item list | BIM ID | Name | Gross quantity | Comment | (Single line edit) | (Single line edit) | (Single line edit) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field) | (Logic field)
PBTR-R | Patient Bay - Resuscitation | FE | FE-11400 | Computer - PC & monitor | 1 | to bench, optional |  |  |  | Yes | Yes | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FE | FE-15500 | Defibrillator | 1 | on resuscitation trolley |  |  |  | Yes | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FE | FE-15900 | Diagnostic set - wall mounted | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FE | FE-3100 | Basin - scrub, Type A | 1 | refer to BHWS-A (sim) |  |  |  | No | No | Yes | No | Yes | Yes | Yes | No | No
PBTR-R | Patient Bay - Resuscitation | FE | FE-32650 | Light - operating, single, 550 diam nominal | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FE | FE-36050 | Monitor - cardiac | 1 |  |  |  |  | Yes | Yes | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FE | FE-3725 | Bed - emergency, electric | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FE | FE-43800 | Refrigerator - under-bench, 120 Litre | 1 |  |  |  |  | Yes | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FE | FE-47000 | Services pendant - ceiling | 1 |  |  |  |  | Yes | Yes | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FE | FE-54510 | Telephone - handset, standard | 1 | to bench |  |  |  | No | Yes | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FE | FE-68800 | Xray viewer - 3 screen | 1 | recessed |  |  |  | Yes | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FE | FE-9800 | Clock - time elapse | 1 | beside analogue |  |  |  | Yes | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-1050 | Bench - laminate | 1 | open under part |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-10700 | IV hanger for ceiling track | 1 | optional, may be attached to pendant |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-11200 | IV track - ceiling mounted | 1 | optional |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-12900 | Wire panel - for storage-bins | 2 | for storage, above bench |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-1470 | Bracket - sharps | 1 | on trolley |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-150 | Air flowmeter | 2 | to services pendant |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-16250 | Oxygen flowmeter | 3 | to services pendant |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-18800 | Trolley - resuscitation, adult | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-19060 | Safe - dangerous drugs, small | 1 | in cupboard |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-1950 | Chair - ergonomic, high | 1 | for staff use |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-19700 | Sharps bin | 1 | on trolley |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-20000 | Shelving unit - adjustable | 3 | over bench, storage |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-21755 | Storage bins | 3 | for storage on louvred panel |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-21950 | Suction adapter | 3 | to services pendant |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-22000 | Suction bottle | 3 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-4375 | Cupboard - medication, lockable | 1 | with drug safe |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-4410 | Cupboard - under bench | 2 | adjustable shelf |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-4467 | Cupboard - wall mounted | 2 | over bench, drug storage |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-4600 | Curtain track - bed screen | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-5000 | Curtain - bed screen | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-5250 | Curtain - privacy screen | 2 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-6152 | Dispenser - disposable glove | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-6154 | Dispenser - paper/ hand towel | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-6156 | Dispenser - soap | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-6158 | Dispenser - soap antiseptic | 1 | to basin |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-5824 | Waste bin - general | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | FF | FF-5975 | Waste bin - clinical | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-9000 | GPO: Single | 2 | fridge, diagnostic set |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-9010 | GPO: Emergency power, single (V) | 16 | services pendant (2-UPS) |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-9015 | GPO: Cleaner | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-9030 | Cardiac protected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-9011 | GPO: Emergency power, double (H) | 1 | for resus trolley |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-9011 | GPO: Emergency power, double (H) | 2 | to bench |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-9060 | GPO: Special: for equipment | 3 | theatre light, clocks, XRV |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-1006 | Voice / Data outlet: double | 2 | 1-services pendant; 1-bench; 1 outlet on each panel for telephone |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-1120 | Monitor connection facilities and cabling | 1 | to services pendant, optional |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-5000 | Airconditioning | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-6001 | General: colour corrected | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-6200 | Light switches | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-6055 | Special: task light, built in | 1 | under O/H shelves/ cupboards |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-8000 | Patient/ Staff call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-8005 | Patient call handset | 1 | call button on the handset |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-8010 | Staff/ Nurse assist call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-8080 | Emergency call | 1 | + indicator button & light |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-1951 | Staff assist call | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-2072 | Scavenge Outlet | 1 |  |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-7000 | Oxygen (O2) | 3 | services pendant |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-7010 | Nitrous Oxide (N2O) | 1 | optional |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-7020 | Medical Air (MA) | 2 | services pendant |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-7040 | Suction | 3 | services pendant |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | SE | S-7070 | Scavenge | 1 | if N2O provided |  |  |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | Fabrics | 1000 | Floor finish | 1 |  | Vinyl | Standard slip resistant | Seamless, coved | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | Fabrics | 3000 | Wall finish | 1 |  | Paint | Acrylic, washable |  | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | Fabrics | 3060 | Wall finish | 1 | to basin & above bench | Vinyl | Prefinished | Splashback | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | Fabrics | 4020 | Wall protection | 1 | As required, equal to 'Acrovyn' | Composite | Prefinished PVC | Crash rail at 900 AFFL | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | Fabrics | 4031 | Wall protection | 1 | As required, Equal to 'Acrovyn' | Composite | Prefinished PVC | Corner guards to 1500 AFFL | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | Fabrics | 5010 | Ceiling - Plasterboard | 1 |  | Plasterboard | Paint, washable | Flush set, suspended | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | Fabrics | 6030 | Cornice | 1 |  | Aluminium | Powdercoat | 24mm shadow angle | No | No | No | No | No | No | No | No | No
PBTR-R | Patient Bay - Resuscitation | Fabrics | 2000 | Skirting | 1 |  | Vinyl | Prefinished | Floor vinyl coved, 150 mmH | No | No | No | No | No | No | No | No | No