## Room Data
Standard Components - Room Data Sheet
Date of issue: 18.01.2021
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
PLST | Plaster Room | 18 | 24 Hours | 1 patient; 2 -4 staff |  | Sound attenuation level - high; acoustic privacy is required | The Plaster Room allows for the application of plaster or other splints and for the closed reduction of displaced fractures or dislocations under sedative or regional anaesthesia. Typically provided as a dedicated room for high volume services, otherwise, a generic Treatment Room may be used with the inclusion of a plaster trap and other specific plastering equipment. The cast saw shall incorporate an integrated dust extraction system or the room shall have an independent dust extraction system suitably located to capture the dust at source. A room exhaust shall also be included to prevent dust from entering adjacent spaces. The plaster trap shall be easily accessible for cleaning and maintenance.

## Items
Standard Components - Items in room
Date of issue: 18.01.2021
Generated by dRofus (C) dRofus
Items
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment
PLST | Plaster Room | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 | floor finish must be minimum slip rating R10; Pendulum P3 or agreed equivalent to extend under basin or inset section of non-slip vinyl required
PLST | Plaster Room | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1
PLST | Plaster Room | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | optional
PLST | Plaster Room | FA | DWPR-005 | DOOR PROTECTION: protection plate to 900 AFFL | 2
PLST | Plaster Room | FA | WLWA-007 | WALL FINISH: splashback | 1 | to basin
PLST | Plaster Room | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1
PLST | Plaster Room | FA | WLWA-020 | WALL FINISH: vinyl, to 1500 AFFL | 1
PLST | Plaster Room | FA | CLFS-006 | CEILING: plasterboard, flush set, suspended, paint, washable | 1
PLST | Plaster Room | FA | DWGL-033 | DOOR: 1400mm c/o, 1 1/2 leaf, observation panel, solid core/glass, paint | 1
PLST | Plaster Room | FA | CLCN-008 | CORNICE: square set flush between wall and ceiling, paint | 1
PLST | Plaster Room | FFE | MGFP-002 | FLOWMETER: medical air | 1
PLST | Plaster Room | FFE | FIHR-019 | BRACKET: for suction bottle | 2 | quantity to suit number of suction outlets
PLST | Plaster Room | FFE | FIDI-001 | DISPENSER: alcohol based hand rub | 1
PLST | Plaster Room | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 3
PLST | Plaster Room | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 3
PLST | Plaster Room | FFE | FQGE-010 | CURTAIN: privacy screen | 1
PLST | Plaster Room | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin
PLST | Plaster Room | FFE | FIDI-014 | DISPENSER: disposable glove | 1 | to basin
PLST | Plaster Room | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin
PLST | Plaster Room | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin
PLST | Plaster Room | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin
PLST | Plaster Room | FFE | MMBE-043 | TROLLEY: patient, examination | 1 | or plinth/examination couch
PLST | Plaster Room | FFE | MGFP-004 | FLOWMETER: oxygen | 2 | quantity to suit number of oxygen outlets
PLST | Plaster Room | FFE | MMMA-011 | MIXER: oxygen and nitrous oxide | 1 | optional, may be mobile, on trolley or wall mounted
PLST | Plaster Room | FFE | MMAH-028 | TROLLEY: plaster | 1
PLST | Plaster Room | FFE | MMGE-002 | ADAPTER: suction | 2 | quantity to suit number of suction outlets
PLST | Plaster Room | FFE | MMGE-100 | SUCTION BOTTLE | 2 | quantity to suit number of suction outlets
PLST | Plaster Room | FFE | HYBA-010 | BASIN: handwash, Type B | 1
PLST | Plaster Room | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin
PLST | Plaster Room | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | to sink
PLST | Plaster Room | FFE | ITNE-008 | COMPUTER: for workstation on wheels | 1 | computer may alternatively be on benchtop or wall mounted; allocation of WOWs across department to be rationalised
PLST | Plaster Room | FFE | ITCL-014 | CLOCK: analogue, electric | 1
PLST | Plaster Room | FFE | MMGE-031 | DIAGNOSTIC SET: wall mounted | 1 | optional, may be portable/mobile on stand
PLST | Plaster Room | FFE | LIFX-027 | LIGHT: examination, ceiling mounted | 1
PLST | Plaster Room | FFE | HYDR-008 | TRAP: plaster | 1 | to be easily accessible for cleaning and maintenance
PLST | Plaster Room | FFE | MMAH-023 | SAW: plaster | 1 | integral/localised dust extraction required
PLST | Plaster Room | FFE | HYBA-053 | SINK: single, 450 x 350 x 300D | 1 | hot water
PLST | Plaster Room | FFE | ITNE-053 | TELEPHONE: handset, wall mount, handsfree | 1
PLST | Plaster Room | FFE | FIJO-060 | BENCH TOP: 600D, joinery under, stainless steel lipped | 1 | with sink and wet edge
PLST | Plaster Room | FFE | FQCW-008 | BIN: waste, clinical, 50L | 1 | optional; type, size and quantity of bins to be determined based on service requirements and waste management policies
PLST | Plaster Room | FFE | MMGE-164 | WORKSTATION ON WHEELS: clinical, mobile | 1 | computer may alternatively be on benchtop or wall mounted; allocation of WOWs across department to be rationalised
PLST | Plaster Room | FFE | MMGE-001 | ADAPTER: scavenge | 1
PLST | Plaster Room | FFE | FIJO-212 | BULKHEAD: joinery | 1 | plasterboard bulkhead also acceptable
PLST | Plaster Room | FFE | FIRT-039 | TRACK: curtain, privacy, Z shaped, angled | 1
PLST | Plaster Room | FFE | MMAH-045 | VACUUM: plaster saw | 1 | or independent dust extraction system
PLST | Plaster Room | SE | ITIN-016 | OUTLET: data, single RJ45, wall mounted | 1 | to telephone
PLST | Plaster Room | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 1
PLST | Plaster Room | SE | ELPR-006 | RCD: residual current device | 4
PLST | Plaster Room | SE | ELGP-101 | GPO: single, ceiling mounted | 1 | to exam light
PLST | Plaster Room | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | + indicator button & light
PLST | Plaster Room | SE | MGAS-053 | OUTLET: scavenge, on services panel | 1
PLST | Plaster Room | SE | LIFX-008 | LIGHT: task light, built in | 3
PLST | Plaster Room | SE | ELBO-015 | SWITCH: light, single | 3 | 1 x general room lighting, 1 x task light, 1 x examination light
PLST | Plaster Room | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 2 | 1 is optional, quantity dependent on service requirements
PLST | Plaster Room | SE | MGAS-032 | OUTLET: nitrous oxide (N2O), on services panel | 1
PLST | Plaster Room | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1
PLST | Plaster Room | SE | MGAS-062 | OUTLET: suction, on services panel | 2 | 1 is optional, quantity dependent on service requirements
PLST | Plaster Room | SE | ITCL-006 | BUTTON: nurse call, patient to staff | 1 | + indicator button & light
PLST | Plaster Room | SE | ITCL-028 | HANDSET: nurse call, with patient to staff call button | 1 | call button on the handset
PLST | Plaster Room | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | + indicator button & light
PLST | Plaster Room | SE | ELGP-109 | GPO: single, wall mounted | 1 | low level, for patient trolley
PLST | Plaster Room | SE | ELGP-208 | GPO: double, wall mounted | 4
PLST | Plaster Room | SE | ELGP-110 | GPO: single, wall mounted, cleaner | 1
PLST | Plaster Room | SE | ITSE-006 | BUTTON: duress, fixed | 1
PLST | Plaster Room | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1
PLST | Plaster Room | SE | ITIN-025 | OUTLET: data, double RJ45, on services panel | 1
PLST | Plaster Room | SE | ELGP-105 | GPO: single, on services panel | 3
PLST | Plaster Room | SE | ELGP-122 | GPO: emergency power, single, on services panel | 3