## RDS
Standard Components - Room Data Sheet
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
AusHFG Room Number | Room name | Prog Area | Hours of Operation | Occupancy | Remarks | Acoustics | Description / Special Requirements
SCRN | Screening Room (Fluoroscopy) | 36 | Available 24 Hours | 1 patient; 1 - 2 staff |  |  | Screening Room (Fluoroscopy) is a room for undertaking radiographic procedures involving contrast media. Patient Change, Toilets and Shower should be located in close proximity. The room may include a preparation area with sink. Additional Design Requirements: Radiation shielding to walls and glazing as advised by Radiation Consultant. Unistrut type ceiling may be required, depending on equipment to be installed. Cable ducting between the X-ray table, vertical bucky and control area may be required, depending on equipment type and layout. 'X-ray in Use' warning lights will be required outside each entry door. Dimmable downlights are optional.

## Items in room
Standard Components - Items in room
Date of issue: 20.03.2017
Generated by dRofus (C) dRofus
Items | [6 empty] | Service allocations
AusHFG Room Code | Room Name | Item list | Item Number | Name | Gross quantity | Comment | Electrical | Data | Cold Water | Hot Water | Warm Water | Tap | Drainage | Gas | Steam
SCRN | Screening Room (Fluoroscopy) | FA | CLCN-002 | CORNICE: aluminium, 24mm shadow angle, powdercoat | 1 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FA | CLTI-008 | CEILING: plasterboard, drop-in tiles, prefinished, 600 x 1200 | 1 | at 3000 mm AFFL | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FA | DWOB-009 | OBSERVATION: internal window, lead glass, clear | 1 | to control booth | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FA | DWPR-006 | DOOR PROTECTION: protection plate, to 900 AFFL, PVC, prefinished | 1 | equal to Acrovyn | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FA | DWPR-010 | DOOR PROTECTION: to door frame, PVC, prefinished | 1 | equal to Acrovyn | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FA | DWSC-003 | DOOR: 910mm c/o, single leaf, solid core, paint | 1 | to Patient Change/ Toilet, number to suit layout | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FA | DWSC-026 | DOOR: 1600mm c/o, 1 1/2 leaf, solid core, lead shielding, paint | 1 | Patient entry/exit, min. 1400mm clear opening | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | ELBO-015 | SWITCH: light, single | 5 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | ELGP-105 | GPO: single, on services panel | 2 | on services panel | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | ELGP-106 | GPO: single, special, for equipment | 2 | Xray Fluoroscopy Equipment, chest bucky | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | ELGP-122 | GPO: emergency power, single, on services panel | 2 | on services panel | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | ELGP-206 | GPO: double, on services panel, RCD protected | 1 | on services panel | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | ELGP-208 | GPO: double, wall mounted | 4 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FIDI-014 | DISPENSER: disposable glove | 1 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FIDI-025 | DISPENSER: paper/hand towel | 1 | to basin | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FIDI-030 | DISPENSER: soap | 1 | to basin | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FIDI-031 | DISPENSER: soap antiseptic | 1 | to basin | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FIHR-013 | BRACKET: for monitor, ECG | 1 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FIHR-044 | RACK: lead apron | 1 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FIJO-020 | BENCH: 600D, open under, laminate finish | 1 | workstation | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FIJO-060 | BENCH TOP: 600D, joinery under, stainless steel lipped | 1 | wet edge | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FIJO-136 | CUPBOARD: underbench, double door, adj shelves, lockable | 2 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FIJO-145 | CUPBOARD: wall mount, double door, adj shelves, lockable | 2 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FA | FLSK-010 | SKIRTING: vinyl, floor vinyl coved, 150H, prefinished | 1 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FA | FLVY-007 | FLOOR FINISH: vinyl, seamless, coved, standard slip resistant | 1 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FQBS-014 | CHAIR: client | 1 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FQCW-016 | BIN: waste, general, 20L | 1 | to basin | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | FQDW-005 | KEYBOARD SHELF: slide out, for workstation | 1 | for PACS computer screen | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | HYBA-010 | BASIN: handwash, Type B | 1 | refer BHWS-B (similar) | No | No | Yes | No | Yes | Yes | Yes | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | HYBA-051 | SINK: single | 1 | warm water | No | No | Yes | Yes | Yes | Yes | Yes | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | HYTP-037 | TAPWARE: mixer, elbow levers | 1 | to basin | No | No | Yes | No | Yes | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | HYTP-038 | TAPWARE: mixer, elbow levers, long | 1 | to sink | No | No | Yes | No | Yes | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | ITCL-004 | BUTTON: nurse call, emergency | 1 | on services panel | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | ITCL-008 | BUTTON: nurse call, staff assist | 1 | on services panel | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | ITIN-026 | OUTLET: data, double RJ45, wall mounted | 3 | 1 outlet for telephone | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | ITNE-009 | COMPUTER: hi performance, for clinical / PACS use | 1 | flat screen | Yes | Yes | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | ITNE-050 | TELEPHONE: handset, desktop, handsfree | 1 |  | No | Yes | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | ITSE-029 | INTERCOM: audio, wall mount | 1 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | MGAS-022 | OUTLET: medical air (MA), on services panel | 1 | on services panel | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | MGAS-042 | OUTLET: oxygen (O2), on services panel | 1 | on services panel | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | MGAS-062 | OUTLET: suction, on services panel | 1 | on services panel | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | MGFP-002 | FLOWMETER: medical air | 1 | to services panel | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | MGFP-004 | FLOWMETER: oxygen | 1 | to services panel | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | MMGE-002 | ADAPTER: suction | 1 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | MMGE-100 | SUCTION BOTTLE | 1 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | MMMI-117 | CONSOLE: control, fluoroscopy/screening | 1 |  | Yes | Yes | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | MMMI-120 | GENERATOR: radiology, fluoroscopy/screening | 1 |  | Yes | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | MMMI-124 | STAND: floor, with patient table, fluoroscopy/screening | 1 |  | Yes | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | MMMI-125 | SYSTEM: imaging, fluoroscopy/screening imaging unit | 1 |  | Yes | Yes | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | MMMI-283 | BUCKY: wall, general X-Ray | 1 |  | Yes | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FFE | MMMI-350 | SCREEN: radiation, lead glass | 1 | control booth | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | MMSP-026 | SERVICES PANEL: wall mounted | 1 | flush mounted | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | SE | WLSH-004 | SHIELDING: lead lining | 1 | to walls and doors as specified | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FA | WLWA-004 | WALL FINISH: paint, acrylic, washable | 1 |  | No | No | No | No | No | No | No | No | No
SCRN | Screening Room (Fluoroscopy) | FA | WLWA-010 | WALL FINISH: splashback, vinyl | 1 | to basin & sink, tiles also acceptable | No | No | No | No | No | No | No | No | No